## Slide 1: LAB MANUAL FOR
COURSE NAME | Mechanical measurement and Instrumentation lab
COURSE CODE | ME 362
CREDIT | 1
ASSESMENT | 50
Mechanical measurement and Instrumentation lab
DEPARTMENT OF MECHANICAL ENGINEERING
COURSE INFORMATION

## Slide 2
Course Learning Outcomes (CLOs) The learning outcomes that are expected to be attained by the student at the end of the course are. |  |  |  | 
Sl. No. | CLOs | Domain of learning | PLOs | 
1 | DISPLAY basic proficiency in operation of the apparatus and PERFORM the experiment to determine the solution of the engineering problems related to the subject. | Psychomotor | 1,2,3 | 
2 | Communicate the learned concepts using different media i.e., verbal and written. As well as perform teamwork. | Affective | 5,6 | 
3 | Manifest the professional responsibilities and norms of engineering practice. | Psychomotor, Affective | 7,10 | 

## Slide 3
CONTENTS
S.No. | NAME OF THE EXPERIMENT | Week
1 | CALIBRATION OF PRESSURE CELL | 1-2
2 | CALIBRATION OF LVDT | 3-4
3 | CALIBRATION OF STRAIN GAUGE | 5-6
4 | ANGULAR DISPLACEMENT BY CAPACITIVE TRANSDUCER | 7-8
5 | SPEED MEASUREMENT | 9-10
6 | CALIBRATION OF McLEOD GAUGE | 11-12
7 | CALIBRATION OF TEMPERATURE TRANSDUCER | 13-14
8 | ROTAMETER SET - UP | 15-16

## Slide 4
AIM:
Calibrate the given Pressure Cell with
Pressure Gauge for its performance.
APPARATUS REQUIRED:
Pressure cell
Hydraulic dead weight Pressure gauge Tester to develop the
pressure Digital pressure indicator Dial type pressure indicator
THEORY:
Pressure is defined as force per unit area and is measured in Newton per square meter (Pascal) or in terms of an equivalent head of some standard liquid.
Thus, a pressure gauge is connected to the hydraulic line and the gauge
1. CALIBRATION OF PRESSURE CELL

## Slide 5
itself stands in atmospheric pressure.
Pressure transducer is basically an electromechanical devices, especially manufactured and designed or wide range application in pressure measurement.
The pressure transducer comprises of diaphragm and an inputs to facilitate pressure measurement. The strain gauges are bonded directly to the sensing member to provide excellent linearity, low hysteresis and repeatability.
Fluid medium whose parameter has to be measured is allowed to deflect the diaphragm (sensing member), which is a single block material and forms an integral part of the pressure transducer.
Usually, the Pressure Transducers are made up non-magnetic stainless steel and thus has the advantage of avoids the yielding effects and leakage problems. The slight deflection of the diaphragms due to the pressure provides an electrical output.
The material most commonly used for manufacture of diaphragms are steel, phosphor bronze, nickel silver and beryllium copper. The deflection

## Slide 6
generally follows a linear variation with the diaphragm thickness.
PANEL DETIALS:
MAINS ON INDICATOR: To indicate the Power given to the system.
CONSOLE ON SWITCH: Provided to activate the system.
PRESSURE INDICATOR: To indicate the Pressure in digital format with Zero knob facility.
SOFTWARE: Facilitates to do things in computer format.
PREPARATION OF EQUIPMENT:
Connect the pressure cell to the pressure indicator with given cable.
Connect the instrument to 1ph, 230V AC supply which is having proper earthing.
Adjust the zero pot of the indicator to indicate zero.
Close the release valve of pressure gauge tester and apply the 10kg dead weight on flange.

## Slide 7
PROCEDURE:
Slowly rotate the screw road in clockwise direction with the help of handle until flange lift up (so that pressure is developed up to applied load). Now observed the digital reading. If it is not showing zero then make it zero by rotating ZERO knob. Now instrument is calibrated.
Apply the load up to 10Kgs one by one on the flange and give pressure by rotating the screw rod such that the dial gauge reads 1 to 10 with respect to load applied.
Note down the readings of dial gauge and pressure indicator, simultaneously in every step.
Calculate the error and % error.

## Slide 8
TABULATIONS:
Sl. No | Pressure in Dial gauge, Pc=kg / cm2 | Pressure in Digital indicator, Pg =kg / cm2 | Correction Pc –Pg | Error Pg –Pc | % Error (Pg – Pc)/ Pg*100
CONCLUSIONS OF THE RESULTS TABULATED:
Summarizing the entire operation
Describing the possible error factors
Techniques which can be adopted to minimize the errors in all aspects i.e., from startup to end.

## Slide 9
LIMITATIONS
Range of Pressure Cell
Max. 10kg/cm^2
APPLICATIONS:
In compressors
In boilers

## Slide 10: 2. CALIBRATION OF LVDT
AIM:
Calibrate Linear Variable Differential Transformer (LVDT) for the performance using Micrometer.
APPARATUS REQUIRED:
LVDT
Digital LVDT indicator
MICROMETER
THEORY BEHIND:
LVDT is an inductive transducer used to translate the linear motion into electrical signal LVDT consists of a single primary winding ‘P’ and two secondary windings (S1 & S2) wounds on a cylindrical armature. An AC source is connected to the primary winding. A movable soft iron core attached with an arm placed inside the armature.

## Slide 11
The primary winding produces an alternating magnetic field which induces alternating voltage in the secondary windings. Single voltage is obtained by connecting the two secondary windings in series. Thus, the output voltage of the transducer is the difference of the two voltages.
When the core is at null position, the flux linking with both the secondary windings is equal. Since both the secondary winding have equal number of turns, M the induced emf is same in them. The output voltage is the difference of the two emf say E1 & E2. When they are equal, the voltage is zero at null position.
When the core is moved to the left side from null position more flux links with S1. The output voltage is V=E2-E1, is greater, the V value is negative (– ve). Means the voltage is read in terms of mm length on the display board indicates the negative value. When the core is moved to the right side of the null position, more flux links with S2 induces voltages which is +ve. The display board indicates the +ve value in mm of length.
The voltage output is linear and is depending on the position of the core. Hence LVDT can be conveniently used to measure the thickness ranging from

## Slide 12: fraction of mm to a few cms. Normally LVDT can give better result up to 5mm.
PANEL DETIALS:
MAINS ON INDICATOR: To indicate the power given to the system.
CONSOLE ON SWITCH: Provided to activate the system.
LVDT INDICATOR: To indicate the Distance moved.
SOFTWARE: FACILITATES TO DO THINGS IN COMPUTER FORMAT.
PREPARATION OF EQUIPMENT:
Make the Micrometer to Read 10mm on the scale. (ZERO POSITION of LVDT)
Connect the instrument to 1ph, 230V AC supply which is having proper earthing.
LIMITATIONS
Range of Pressure LVDT: -10mm to +10mm

## Slide 13: PROCEDURE:
Slowly rotate the screw head of the micrometer either clockwise or anticlockwise to measure 1mm on it.
Note the Reading on the LVDT indicator.
Repeat step 1 and step 2 until 10mm on either side.
Note down the readings of Micrometer, simultaneously in every step.
Calculate the error and % error.
Note: Clockwise will give readings in –ve direction and Anticlockwise will give in +ve direction.

## Slide 14: TABULATIONS:
Sl. no | Actual Reading, ‘Ra’ mm | Measured Reading, ‘Rm’ mm | Error ‘E’ | % Error
RESULT: -----------------------
Note: Plot a graph between Ra Vs Rm
APPLICATIONS:
1. In pressure cells
2. In force cells
3. In accelerometers

## Slide 15: 3. CALIBRATION OF STRAIN GAUGE
AIM:
To determine the elastic constant (modulus of elasticity) of a Cantilever beam subjected to concentrated end load by using STRAIN GUAGES.
APPARATUS REQUIRED :
Load Cell with Strain Gauge Digital
Strain indicator Weights
THEORY BEHIND:
A body subjected to external forces is in a condition of both stress and strain. Stress can be directly measured but its effect i.e., change of shape of the body can be measured. If there is a relationship between stress and strain, stresses occurring in a body can be computed if sufficient strain information is available. The constant connecting the stress and strain in elastic material under the direct stresses is the modulus of elasticity,

## Slide 16
i.e., E=σ / є
the principle of the electrical resistance strain gauge was discovered by Lord Kelvin, when he observed that a stress applied to a metal wire, besides changing resistance strain gauges are made into two basic forms, bonded wire and bonded foil. Wire gauges are sandwiched between two sheets thin paper and foil gauges are sandwiched between two thin sheets of epoxy.
The resistance factor ‘R’ of a metal depends on its electrical resistively, , its area, a and the length l, according to the equation R = l / a.
Thus, to obtain a high resistance gauge occupying a small area, the metal chosen has a high resistively, a large number of grid loops and a very small cross-sectional area. The most common material for strain gauge is a copper - - nickel alloy known as Advance. The strain gauge is connected to the material in which it is required to measure the strain, with a thin coat of adhesive. Most common adhesive used is Eastman, Deco Cement, etc. as the test specimens extends or contracts under stress in the direction of windings, the length and cross-sectional area of the conductor alter, resulting in a corresponding increase or decrease in electrical resistance.

## Slide 17
GAUGE FACTOR:
The dimension less relationship between the change in gauge resistance and change in length is called Gauge factor of the strain, which is expressed mathematically,
Gauge Factor, g = (R/R) / (l/l)
In this relationship R and I represent, respectively the initial resistance and initial length of the strain gauge filament, while R & l represents the small change in resistance and length, which occurs as the gauge is strained along with the surface to which it is bonded. This gauge factor of a gauge is a measure of the amount of resistance change for a given strain. The higher the gauge factor greater the electrical output for indication or recording purpose. The gauge factor is supplied by the manufacturer and may range from 1.7 to 4.
The usual method of measuring the change of resistance in a gauge element is by means of Wheatstone bridge as shown in figure. It consists of Galvanometer, 4 resistor & a battery. Resistance R1 is the strain gauge is used

## Slide 18
for strain measurement, which is mounted on the specimen. The three resisters R2, R3 and R4 are internal to the device.
Let us assume that the resistance has been adjusted so that the bridge is balanced.
The most common bridge arrangements are single arm, two arm and four arm mode.
i.e. Voltage Ebd = 0.

## Slide 19: Single Arm Mode (Quarter bridge).
This bridge arrangement consists of a single active gauge in position, say R1 and three resistor are internal to the device. Temperature compensation is possible only if a self temperature compensating strain gauge is used.
Two Arm Mode (Half bridge).
In this mode, two resistors are internal to the device and the remaining two are strain gauges. One arm of this bridge is commonly labeled as active arm and the other as compensating arm. The bridge is temperature compensated.
Four Arm Mode (Full Bridge).
In this bridge arrangement, four active gauges are placed in the bridge with one gauge in each of the four arms. If the gauges are placed on a beam in bending as shown in fig of the elastic constant by bending test experiment, the single from each of the four gauges will add. This bridge arrangement is temperature compensated.

## Slide 20: PANEL DETIALS:
MAINS ON INDICATOR: To indicate the Power given to the system.
CONSOLE ON SWITCH: Provided to activate the system.
STRAIN INDICATOR: To indicate the Distance moved.
SOFTWARE: FACILITATES TO DO THINGS IN COMPUTER FORMAT.
PREPARATION OF EQUIPMENT:
Connect the instrument to 1ph, 230V AC supply which is having proper earthing.
Switch on the Console
Select the BRIDGE Mode
Connect the Strain wires accordingly in the table below
5. Make the Indicator to Read Zero. (ZERO POSITION)
Prepare the Loads to be added

## Slide 21
Sl. No | Bridge Mode | Color | No of Wires | Connect To
1 | Quarter | ORANGE | 2 | A & B
2 | Half | ORANGE | 2 | A & B
 |  | GREEN | 2 | B & C
3 | FULL | WHITE | 2 | A & D
 |  | ORANGE | 2 | A & B
 |  | GREEN | 2 | B & C
 |  | BLACK | 2 | C & D
LIMITATIONS
Range of Load cell: 10kg

## Slide 22: PROCEDURE:
For the Bridge selected and connected wires, slowly add the Weights in steps of 1kg.
Note the Reading on the Strain indicator
Repeat step 1 and step 2 until 10kg is loaded.
Note down the weights added, simultaneously in every step.
Calculate the error and % error
Modulus of Elasticity of the given load cell.

## Slide 23: TABULATIONS:
Sl. | Load |  | Strain | Measured | Bending | Modulus of
No | Applied |  | Indicator | Strain | stress. | elasticity.
 | W (N) |  | Reading | Єm=є*10 / | σ = 6wl / bh | E = σ / Єm
 |  |  | Є – micro | No. of Bridge |  | (N/nm )
 | W | N |  |  |  | 
 |  |  | strain |  |  | 

## Slide 24: CONCLUSIONS OF THE RESULTS TABULATED:
Summarizing the entire operation
Describing the possible error factors Graph Plotting
Techniques which can be adopted to minimize the errors in all aspects i.e, from startup to end.
APPLICATIONS:
1. In the ropes
2. In in the beams

## Slide 25: 4. MEASUREMENT OF ANGULAR DISPLACEMENT
AIM:
Measurement of angular displacement using capacitive transducer.
APPARATUS REQUIRED:
CAPACITIVE ANGULAR DISPLACEMENT CELL
Digital ANGULAR indicator
STEPPER MOTOR ASSEMBLY with Controller
THEORY BEHIND:
Capacitance is well known to be a function of effective area of the conductors, separation between them, the dielectric strength of the material in the separation. Capacitive transducers convert the physical quantity to be measured into a change of capacitance which is processed by the measuring circuit of the transducer. The capacitance of a parallel plate capacitor may be

## Slide 26: changed by varying the separation between the plates, varying the effective area of the plates or varying the dielectric.
Capacitive type transducers are used essentially for displacement or positioning measurements. But they are more susceptible to environmental factors such as dust or moisture in the atmosphere than inductive type transducers.
The meshing area between two stator and rotor plates of the capacitor goes on changing as the shaft capacitor is rotated. The arrangement is used to demonstrate the measurement of angular displacement. The transducer is mounted on to the face of a protractor which indicates the angle of displacement, and the readout display the amount of displacement.
PANEL DETIALS:
MAINS ON INDICATOR: To indicate the Power given to the system.
CONSOLE ON SWITCH: Provided to activate the system.

## Slide 27: ANGULAR INDICATOR: To indicate the Distance moved.
STARTER:
To Start the Motor and to Select the Mode of Operation
INCREMENTAL:
To Give Step input to the Motor according to mode selected.
SOFTWARE:
FACILITATES TO DO THINGS IN COMPUTER FORMAT.
PREPARATION OF EQUIPMENT:
Connect the instrument to 1ph, 230V AC supply which is having proper earthing
Press the STARTER switch and select MANUAL MODE.
Set ZERO position on the Manual Indicator by pressing the INCREMENTAL Button.

## Slide 28: LIMITATIONS
Range of ANGULAR DISPLACEMENT: Up to 360deg
In AUTO Mode the Incremental will be in steps of 22.5deg.
PROCEDURE:
  After the Preparation of equipment, Now press the INCREMENTAL button so the stepper motor sets accordingly with the selected mode of previously.
  Note the Reading on the Angular Displacement indicator
  Repeat step 1 and step 2 until required.
  Note down the readings of Manual Angular, simultaneously in every step.
  Calculate the error and % error.
Note: Rotation will be in Clockwise direction only.

## Slide 29: TABULATIONS:
Sl. no | The displacement protractor reading | Measured displacement meter reading | Deviation
CONCLUSIONS OF THE RESULTS TABULATED:
Summarizing the entire operation
Describing the possible error factors Graph Plotting
Techniques which can be adopted to minimize the errors in all aspects i.e., from startup to end.
APPLICATIONS:
1. In the rotating shafts

## Slide 30: 5. CALIBRATION OF PHOTO SPEED SENSOR
AIM:
Calibrate PHOTO SPEED SENSOR for the performance using MAGNETIC Speed Sensor.
APPARATUS REQUIRED:
PHOTO & MAGNETIC SPEED SENSORS
Digital SPEED indicators MOTOR
with Controller.
THEORY BEHIND:
The measurement of rotational velocity is more common. For velocity (speed) measurement the most convenient calibrator scheme uses a combination of toothed wheel, a simple magnetic proximity pickup a photo couple sensor and an electronic indicator to measure the speed. The angular rotation is provided by some adjustable speed drive of adequate stability. The toothed wheel

## Slide 31: mounted with iron rods while passing under magnetic and photo pickup produces an electric pulse. These pulses are fed to signal conditioner unit and displays reading visually. The stability of the rotational drive is easily checked by observing the variation of display reading. PANEL DETIALS:
MAINS ON INDICATOR: To indicate the Power given to the system.
CONSOLE ON SWITCH: Provided to activate the system.
SPEED INDICATORS: To indicate the SPEED.
MOTOR CONTROLLER: To run the motor at various speeds.
SOFTWARE: FACILITATES TO DO THINGS IN COMPUTER FORMAT.
PREPARATION OF EQUIPMENT:
Keep the Speed controller knob in minimum position.
Connect the instrument to 1ph, 230V AC supply which is having proper earthing.
Switch on the Console to activate the panel.

## Slide 32: LIMITATIONS
Maximum RPM: 2500rpm
PROCEDURE:
Once the equipment is prepared, slowly rotate the Speed controller to set the speed desired in range.
Note the Reading on the PHOTO SENSOR PICKUP Indicator
Repeat step 1 and step 2 until required within the range.
Note down the readings of MAGNETIC PICKUP, simultaneously in every step.
Calculate the error and % error.
Note: Clockwise to increase the speed and Anticlockwise to reduce the speed.

## Slide 33
Sl. no | Actual Reading, ‘Ra’ mm | Measured Reading, ‘Rm’ mm | Error E = (Ra- Rm) | % Error {(Ra- Rm)/Ra}*100
TABULATIONS:
Note: Plot a graph between Ra Vs Rm
APPLICATIONS:
To measure the electric motors shaft speed.
To measure the turbine shaft speed.
To measure the engines shaft speed.
RESULT: -------------------

## Slide 34: 6. CALIBRATION OF McLeod GAUGE
AIM:
Calibrate the Vacuum Cell (McLeod Gauge) in comparison with Vacuum Dial Gauge.
APPARATUS REQUIRED :
CYLINDRICAL CELL
Digital Vacuum Gauge Analog Vacuum Gauge
High-Capacity Vacuum Pump.
THEORY BEHIND:
In everyday usage, vacuum is a volume of space that is essentially empty of matter, such that its gaseous pressure is much less than atmospheric pressure. The word comes from the Latin for “empty”. A perfect vacuum would be one with no particles in it at all, which is impossible to achieve in practice. Physicists often discuss ideal test results that would occur in a perfect

## Slide 35
vacuum, which they simply call “vacuum” or “free space”, and use the term partial vacuum to refer to real vacuum. The Latin term in vacuum is also used to describe an object as being in what would otherwise be a vacuum.
Vacuum is useful in a variety of processes and devices. Its first widespread use was in the incandescent light bulb to protect the filament from chemical degradation. The chemical inertness produced by a vacuum is also useful for electron beam welding, cold welding, vacuum packing and vacuum frying. Ultra-high vacuum is used in the study of atomically clean substrates, as only a very good vacuum preserves atomic-scale clean surfaces for a reasonably long time (on the order of minutes to days). High to ultra-high vacuum removes the obstruction of air, allowing particle beams to deposit or remove materials without contamination. This is the principle behind chemical vapor deposition, physical vapor deposition, and dry etching which are essential to the fabrication of semi conductors and optical coatings, and to surface science. The reduction of convection provides the thermal insulation of thermos bottles. Deep vacuum lowers the boiling point of liquids and promotes low temperature out gassing which is used in freeze drying, adhesive preparation, distillation, metallurgy, and process purging. The electrical properties of

## Slide 36: vacuum make electron microscopes and vacuum tubes possible, including cathode ray tubes. The elimination of air friction is useful for flywheel energy storage and ultracentrifuges.
PANEL DETIALS:
MAINS ON INDICATOR: To indicate the Power given to the system.
CONSOLE ON SWITCH: Provided to activate the system.
VACUUM ON SWITCH: Provided to activate the system.
VACUUM INDICATORS: To indicate the Vacuum inside the cylinder.
PREPARATION OF EQUIPMENT:
Make sure the Vacuum On switch is in off condition.
Connect the instrument to 1ph, 230V AC supply which is having proper earthing.

## Slide 37: LIMITATIONS
Maximum Vacuum Range: 600mm of Hg or 800mbar
PROCEDURE:
Switch on the Console to activate the Panel.
Close the Valve on the Cylinder.
Start the Vacuum pump and allow attaining maximum vacuum.
Note down the Readings on Digital and Analog gauges.
Now slowly open the Valve on the cylinder and set to the required Vacuum.
Once again note down the Digital and Analog readings.
Repeat Steps 5 and 6 until zero.
Calculate the error and % error.
TABULATIONS:

## Slide 38
Sl. no | Actual Reading, ‘Ra’ mm | Measured Reading, ‘Rm’ mm | Error E = (Ra- Rm) | % Error {(Ra- Rm)/Ra}*100
TABULATIONS:
Note: Plot a graph between Ra Vs Rm
APPLICATIONS:
To clear halls
RESULT: -------------------

## Slide 39: 7. CALIBRATION OF TEMPERATURE SENSOR
AIM:
Calibrate TEMPERATURE SENSORS for the performance using STANDARD water bath.
APPARATUS REQUIRED:
STANDARD WATER BATH with Temperature Controller
Temperature Sensors like RTD, J – Type, K – Type & T – type Individual Digital Temperature Indicators
THEORY BEHIND:
RESISTANCE TEMPERATURE DETECTOR (RTD)
Resistance thermometers, also called resistance temperature detectors

## Slide 40
or resistive thermal devices (RTD), are temperature sensors that exploit the predictable change in electrical resistance of some materials with changing temperature. As they are almost invariably made of platinum, they are often called platinum resistance thermometers (PTR). They are slowly replacing the use of thermocouples in many industrial applications below 600oC, due to higher accuracy and repeatability.
There are many categories like carbon resistors, film and wire wound types are the most widely used.
Carbon resistors are widely available and are very inexpensive. They have very reproducible results at low temperatures. They are the most reliable from at extremely low temperatures. They generally do not suffer from significant hysteresis or strain gauge effects.
Film thermometer have a layer of platinum on a substrate, the layer may be extremely thin, perhaps one micrometer. Advantages of this type are relatively low cost and fast response. Such devices have improved performance although the different expansion rates of the substrate and platinum give “strain gauge” effects and stability problems.

## Slide 41
Wire – wound thermometers can have greater accuracy, especially for wide temperature ranges.
Coil element has largely replaced wire wound elements in industry. This design has a wire coil which can expand freely over temperature, held in place by some mechanical support which lets the coil keep its shape.
THERMOCOUPLES
The common electrical method of temperature measurement uses the thermocouple, when two dissimilar metal wires are joined at both ends, an emf will exist between the two junctions, if the two junctions are at different temperatures. This phenomenon is called Setback effect. If the temperature of one junction is known then the temperature of the other junction may be easily calculated using the thermoelectric properties of the materials. The known temperature is called reference temperature and is usually the temperature of ice. Potential (emf) is also obtained if a temperature gradient along the metal wires. This is called Thomson effect and is generally neglected in the temperature measuring process. If two materials are

## Slide 42
connected to an external circuit in such a way that current is drawn, an emf will be produced. This is called as Peltier effect. In temperature measurement, setback emf is of prime concern since it is dependent on junction temperature.
The thermocouple material must be homogeneous. A list of common Thermocouple materials in decreasing order of emf chrome, iron and copper platinum – 10% rhodium, platinum, alumel and constantan (60% copper and 40% nickel). Each material is thermoelectrically positive with respect to the below it and negatives with respect those above.
The material used in the Thermocouple probe is:
Iron – Constantan (Type J)
Copper – Constantan (Type T)
Chromyl – Alumel (Type K)

## Slide 43: PANEL DETIALS:
MAINS ON INDICATOR: To indicate the Power given to the system.
CONSOLE ON SWITCH: Provided to activate the system.
TEMPERATURE INDICATORS: To indicate the Temperatures of various sensors.
WATER BATH SETTINGS: To Set the Temperature of Water bath.
SOFTWARE: FACILITATES TO DO THINGS IN COMPUTER FORMAT.
PREPARATION OF EQUIPMENT:
Fill the water into the Water Bath to the required level.
Make sure the water bath is in off condition.
Connect the instrument to 1ph, 230V AC supply which is having proper earthing.
LIMITATIONS
Maximum Temperature: up to 100deg

## Slide 44: PROCEDURE:
Switch on the Console to activate the Panel and then switch on the water bath.
Set the required Temperature of the bath
Select the Sensor under study {however other sensors will also be working condition, this is only done to concentrate on particular sensor type}and Note the readings of the selected Sensor indicator for every two degree rise of water bath temperature.
Repeat step 2 and step 3 until required within the range.
Repeat the above for different sensors.
Calculate the error and % error.

## Slide 45
Sl. no | Actual Reading, ‘Ra’ mm | Measured Reading, ‘Rm’ mm | Error E = (Ra- Rm) | % Error {(Ra- Rm)/Ra}*100
TABULATIONS:
Note: Plot a graph between Ra Vs Rm
APPLICATIONS:
1. To control of gas flow
2. In electric iron boxes
3. In domestic ovens
RESULT: -------------------

## Slide 46: 8. ROTAMETER SETUP
AIM:
The experiment is conducted to know how to Calibrate Rotameter at different flow rate.
PROCEDURE:
Fill in the sump tank with clean water.
Keep the delivery valve closed.
Paste the Log sheet from zero marking on the rotameter to its full height and make marking if necessary.
Connect the power cable to 1Ph, 220V, 10 Amps with earth connection.
Switch on the pump & open the delivery valve.
Adjust the flow through the control valve of the pump.

## Slide 47
Set the height from the log sheet sticked on the rotameter and note the height in cm the rotameter reading in lpm.
Note down the differential head reading in the Manometer. (Expel if any
air is the by opening the drain cocks provided with the Manometer.)
Operate the Butterfly valve to note down the collecting tank reading against the known time and keep it open when the readings are not taken.
Change the flow rate and repeat the experiment.
OBSERVATIONS:
Sl. No | Time for ‘R’ cm rise in water ‘T’ sec | Rotameter Reading, ‘Q’ Lpm | Height measured on Log scale, cm
1 |  |  | 
2 |  |  | 
3 |  |  | 
4 |  |  | 

## Slide 48
TABULAR COLUMNS AND GRAPHS:
For Rotameter:
Height measured on Log scale, cm | Actual Discharge QA m³/sec | Theoretical Discharge QTR m³/sec | Coefficient of Discharge ‘CD’ | Average ‘CD’

## Slide 49: Graphs
Draw graph of Actual discharge Vs Height
on Log scale Draw graph of Theoretical Discharge Vs Height on log scale
RESULTS:
For Rotameter:
1. Actual Discharge, QA | = | m³/s
2. Theoretical Discharge, QT | = | m³/s
3. Co – efficient of Discharge, | Cd | =

## Slide 50: PRECAUTIONS
1) Do not run the pump dry.
2) Clean the tanks regularly, say for every 15days.
3) Do not run the equipment if the voltage is below 180V.
4) Check all the electrical connections before running.
5) Before starting and after finishing the
experiment the main control valve should be in close position.
6) Do not attempt to alter the equipment as this may cause damage to the whole system.